DBBC.2 MANUAL

ver.2.3

INDEX

Introduction

1. System Overview Description

1.1 Description

2. Mechanical Set

2.1 Description

3. Cooling Set

3.1 Description

3.2 Maintenance

4. Power Distributor

4.1 Description

4.2 Maintenance

5. PC Set

5.1 Description

5.2 Maintenance

6. Conditioning Module

6.1 Description

6.2 Maintenance and calibration

7. CaT Set

7.1 Description

7.2 Maintenance and calibration

8. FiLa Boards

8.1 Description

8.2 Maintenance and calibration

9. ADB1 Board

9.1 Description

9.2 Maintenance and calibration

10. Core2 Board

10.1 Description

10.2 Maintenance and calibration

11. Connection Set

11.1 Description

11.2 Maintenance and calibration

12. Firmware Set

12.1 Description

12.2 List of configurations files

13. Software Set

13.1 Description

14. Command Set

14.1 Description

14.2 List of commands

15. Schematics Drawing

15.1 Connection Set

15.2 Power Distributor

15.3 Conditioning Module

15.4 CaT Set

15.5 FiLa

15.6 ADB1

15.7 Core2

16. Firmware _4ch register description (dbbc2.bit)

Introduction

This document describes the DBBC system in the 2009 implementation. The main component parts are described in the functionality in order to allow the customer how to handle maintenance and troubleshooting of the system. No distinction is done between User Manual and System Manual and both types of information are reported embedded one to the other, leaving the relevant elements be hopefully more straightforward to be used.

The DBBC system is a general purpose instrument and it is well adopted as VLBI back-end just because during its development a standard VLBI terminal, as diffused around the world , was taken into account as element the new system had to emulate. This functionally alone could not justify the development process as it was expected to achieve more flexible performance. For this main reason the DBBC system should not be considered a copy of a MKIV back-end system and the main distinctions need to be kept in mind for an appropriate use of the very powerful system that allows to add to a radio-telescope equipment a very general purpose data processing element.

1. System Overview Description

A general description of the DBBC system is given in this chapter with the goal to give the customer the possibility to easily recognize the main parts with a distinction for functionality.

A general overview of the system is showed in the figure 1.1. It could be noticed as the different elements are disposed in a way that could simplify the maintenance operations, and help to create its own system is simplified manner.

A distinction has been done in order to define functional blocks. A list follow:

a)  Mechanical Set

It includes all the mechanical parts, that are relevant for the proper functionality of the system.

b)  Cooling Set

It includes the electrical and mechanical elements in charge for a proper cooling of the entire system.

- Power Distributor

All the elements used for powering the DBBC electronics are kept under this definition, including the power supply for the PC Set.

- PC Set

It’s the group of parts that are adopted for managing the system and have a dialog with the external world.

e)  Conditioning Module

The modules are the relevant interface between an analog source, like a receiver, and the system before the analog to digital conversion.

f)  CaT

Two boards are in charge for generate all the relevant signals related to the Clock and Timing functionalities.

g)  FiLa Boards

These boards are the digital edge interfaces between the external side and the DBBC main functional block.

h)  ADB1 Boards

They represent the analog edge interface between the external side and the DBBC main functional block.

- Core1 Boards

The processing elements in the DBBC Block.

- Connection Set

It represent the entire group of cables that are used to route all the signals, in the analog or digital domain between one to the other element of the system.

l)   Firmware Set

The collection of FPGA configuration files to be used in the Core Boards.

m)  Software Set

The collection of processor files able to manage the entire system.

n)  Doc Set

This document is part of the set that includes also all the information reported in the DBBC web site.

ADBoard

min 1- max 4

CoreBoard

min 1 – max 16

VSI

IF2

IF3

IF4

10 MHz

1 PPS

IF1

PC and

Interfaces

FiLa

FiLa

Ethernet

D

I

S

K

JTAG

Adapter

Monitor

IF1

IF2

IF4

FiLa

ConditioningModule

min 1 – max 4

ConditioningModule

FiLa

PowerDistributor

IF3

CaT

Fig.1.1 DBBC top view

2. Mechanical Set

2.1 Description

The DBBC is mechanically assembled in a 19 inch sub-rack box, 8 U height, 500 mm deep, so to include everything is necessary for a complete standard system. In the front panel, lower right side a blue led indicates when the main power switch is activated. In the rear panel mere elements are present for the connection. In the left lower side there are three switches, one big, two. The big one is just close to the mains connection. The 200VAC power cord should be there applied. It should be noticed as in the same block is present also a general protection fuse. The general switch when ‘on’ activate the cooling system in the DBBC mechanical box, giving also access to the further two smaller power switches, one for the general electronics, one for the internal PC system.

It can be noticed as the mechanics adopted is not composed by general purpose mechanical elements, but built from a company under drawing for the developed and dedicated parts. The thickness of the aluminum foils used is also conceived taking into account a robust and functional aspect.

The mechanical assembly is divided in virtually four parts internal to the box that can be considered dividing with two planes crossing one each other, one parallel to the up-down panels and one parallel to front-rear side panels. The figures 2.1 and 2.2 report the different sections.

Electronics/PCSet/CondMod

Cooling/Power

Fig. 2.1. Front view with two sections

Electronics

PCSet/CondMod

Cooling

Power

Fig. 2.2. Side view with four sections

The skeleton supporting the system is composed by the two lateral panels, in conjunction with six large bars joining the two panels. These alone constitute a structure where are mounted all the remaining parts. The lower front part is occupied by the fan closed in a box, supporting the upper electronics elements with a flexible structure in terms of positioning and number of  DBBC boards in blocks. Several bar elements allow to place appropriately the boards still having the possibility to fix them in the needed position. A rear side half deep plane is used for supporting in the upper side the PC Set and Conditioning Modules, while in the lower side the three power suppliers used one for the PC Set, two for the electronics.

Two black panels with serigraphy are used in the front side. Under the upper front panel is mounted a second transparent Plexiglas panel, that can be used removing the first for having the possibility to look at the led indicators, still maintaining the appropriate cooling conditions with the air flow. Six panels are adopted in the rear side for the PC Set external connection, four Conditioning Modules, Power Distributor.

The upper and lower big panels cover the system and can be easily  removed in case of need, because they do not support any other part.

In the fig 2.3 is showed the rear panel with the functional distribution of the modules placed in the rear side.

PC Set

Cond

Mod

Cond

Mod

Cond

Mod

Cond

Mod

Power

Service

Fig. 2.2. Side view with four sections

3. Cooling Set

3.1 Description

The cooling system is an important element because it allows to properly assure a cooling air flow to the entire box, with a flow direction that is mandatory for keeping under proper conditions all the parts of the system. The air enter the system from the lower front panel, is captured by a tangential fan, then pushed to the half up side of the box, just under the functional DBBC boards block, then left to go to the rear side of the box, to a half upper and half lower when cooling is produced for the PC Set and to the Power Distribution Set. The fig 3.1 reports the air flow directions.

Fig. 3.1. The air cooling flow from a side view

3.2  Maintenance

Care should be taken to assure that the air flow is always appropriate and obstruction is in the air path. Periodic check should be made to the front panel thin net internal to the lower panel, and in case of accumulation of durst, it should be removed and cleaned with compress air.

A general removal of the internal dust is recommended once every 6 months with the help of compress air and vacuum cleaner.

4. Power Distributor

4.1 Description

The Power Distributor is the element that starting from the 220VAC mains produce the different voltages necessary for the different parts of the system. It is made by the following elements:

a) two identical units of switching power supply that produce each the power rails:

12V /13 A, +5V /40 A, +3.3V/40A, -12 V / 5A, total power 350W

b) one switching power supply for the PC Set, 350W

c) Power Distributor PCB, placed in the middle of ht box as separation between the analog and digital part, separating the PC Set and Conditioning Modules in one side and the digital electronics on the other side. This board receive the voltages from the switching power suppliers and generates what is necessary for the different boards of the DBBC stack in appropriate connectors placed in the top of the board. In particular in the pcb front panel it will be generated what necessary for:

c.1) one connector for CaT boards,

c.2) one connector for the first FiLa board

c.3) four connectors for ADB1 boards

c.4) sixteen connectors for Core2 boards

c.5) one connector for the last FiLa board

In the rear side of the pcb are present:

c.6) four connectors for powering the Conditioning Modules.

The Power Distributor PCB separates +5, +3.3V with appropriate filtering for the necessary different sections of all the boards and even in the same board for analog and digital sections, or even for different digital sections. Moreover in the pcb are produced such voltages that are not present on the switching power suppliers. This is necessary for the -5.2V used in the ADB1 boards, and for the 1.5V widely used as VCore in the Core1 boards. This last in particular requires large amount of current and a single additional DC/DC converter is adopted for every Core1 board, each having the capability of 7A at 1.5V.

Single separate multi-wire cables are used to connect the Power Distributor PCB with every board in the stack, the CaT boards, and the Conditioning Modules.

In the tables below are reported voltages at the output of the Power Distributor PCB connectors.

|  | Pin1 | Pin2 | Pin3 | Pin4 | Pin5 | Pin6 |
| --- | --- | --- | --- | --- | --- | --- |
| CaT | +3.3 | +5.0 | -5.2 | +3.3 | +12 | GND |
| FiLa f/l | +3.3 | GND | +3.3 | GND | +3.3 | GND |
| ADB1 | +3.3 | +5.0 | +5.0 | -5.2 | GND | GND |
| Core2 | +3.3 | +3.3 | +3.3 | GND | GND | GND |
| CondMod | +12 | -5.2 | +5.0 | +3.3 | GND | GND |

4.2 Maintenance

Check for proper voltage values once every year the output Power Distributor PCB connectors. Values are regular if in a range of  +- 2% with respect to the nominal value.

5. PC Set

5.1 Description

The DBBC system has an internal industrial PC that is adopted for supporting a number of functionalities. In order to accomplish these tasks a single board PC is used on a small PCI bus together with three additional elements. Two of them are commercial boards from ADLink named PCI 7200 and PCI9111HR, the third is still a commercial element from Xilinx acting as JTAG interface.

Let see the list of tasks supported by the PC Set:

a) download of the Firmware Set loaded from the system disk (hard disk or flash) to the Core1 boards.

b) management of the settings of the Core1 firmware configurations and readings of the parameters.

c) program of the CaT synthesizer for the appropriate 1024 MHz system clock generation.

d) management of the settings of the Conditioning Modules and readings of parameters.

e)  management of observative DBBC schedule.

f)  Field System interface.

The single board PC is Pentium 4 based, operating at 3.0GHz, with a support of  1GB RAM. The operating system is Windows XP because not all the drivers of the used devices are available under a single Linux version or not available at all for Linux. The system disk is a 80GB hard disk, but in some systems is available a disk 4 or 8 GB flash memory.  In the rear side of the system connections are possible for a keyboard, a mouse, a USB2 port, a monitor. A wireless keyboard and mouse is furnished with the system, operating at low frequency. It is responsibility of the user to decide to adopt or not such device for possible RFI reasons.

The PCI 7200 commercial board is able to guarantee IO at 32 bit on separated bus, so two cables connect such board with the first FiLa board, digital interface accessing point to the DBBC stack. All the dialogue, register setting, register reading with the DBBC Stack is realized through this channel.

The PCI9111HR commercial board is used to drive the Conditioning Modules. It is able to read a number of analog channels with 16-bit resolution and for such functionality adopted to read the IF Total Power signal as coming form the Conditioning Module in the selected total band. Moreover the Conditioning Modules require some register setting and for such reason is still used a digital interface bus in the PCI9111HR. A single flat cable is connecting the PCI board with the top side of the Conditioning Modules providing the interface communication channel.

The Xilinx download cable is used to quickly load the configuration files to the Core1 chain, still maintaining the possibility to access the single board, as a JTAG scheme and interface is adopted for such functionality. The connection speed is fixed at 12 Mbps, but could be increased up to 30Mbps. A wide level of flexibility is assured by this Boundary Scan connection because a number of internal tests can be realized in the single board where this could be worth to be done.

5.2 Maintenance

It is recommended to clean from the dust the processor fan and keep the processor area free from any object that could affect the correct cooling of the processor. The PC Set is mounted on a single mechanical block, so it can be easily removed from the DBBC box for cleaning or check reasons. The access is from the rear side of the system.

6. Conditioning Module

6.1 Description

The punctionality of the system is the following:

a)  pre-AD conversion amplitude signal conditioning

b)  pre AD conversion Nyquist band definition

c)  4 RF input selection

d)  output level control

e)  Total Power measurement control

f)  2.2 GHz bandwidth

The connection between the IF signals coming from the receiver to the DBBC system is realized by the Conditioning Module. In the rear side of the DBBC are placed four separated modules, one for each IF. Two input SMA connectors are present in the Conditioning Module and can be used for two different sources with a software selectable switch.

The selected signal is then routed after level adjustment to a set of filters, one from 10 to 512 MHz, the other from 512 to 1024 MHz. The edge frequencies are defined at -15 dB. A switch select the needed frequency band and finally the signal is routed to the output and to the total power measurement detector. Such circuitry produce a voltage level proportional to the power in the band and is available to the rear part of the module together with the signal to be routed to the ADB1 boards.

The system having a programmable attenuator and a total power measurement can realize a closed loop level control with the closure realized by a control software. Indeed the total power level is sent to the PCI9111HR, making part of the PC Set, where is converted in the digital domain at 16-bit and kept available for the IF control software.

The system bandwidth is form 10MHz up to 2.2GHz, and the maximum number of filters selectable by the control software is four.

6.2 Maintenance and calibration

The Conditioning Module doesn’t need any particular maintenance, and the calibration of the Total Power measurement system is not normally required, with the exception of  a post-repairing event.

7. CaT Set

7.1 Description

The Clock and Timing Set represent the elements generating the necessary clock and timing signals from the external reference clock and time data coming from an H-Maser. The system clock used for sampling and for generate all the further clock signals is a 1.024GHz, that is produced by multiplication from the reference 10MHz coming from the rear external Service panel. Such panel reports all the in and out signals in support of the system.

The Clock board of the CaT receive the external reference 10MHz at a nominal level of 0 dBm and reproduce a copy available in the Service panel. A pretty flexible PLL closure scheme is adopted to generate the 1.024MHz necessary as system clock. Indeed a VCO is phase locked in loop in a programmable fractional divider scheme in order to get the possibility to generate a wide number of possible useful clock values (es. 2^31Hz) with very good phase noise performance. Six single-ended SMA output connectors are available on the top of the board, the first four used for driving the four ADB1 boards, the second two used to drive the Timing board and a monitor connector in the Service panel. The board is mounted on a dedicated mechanical assembly to be fixed in the general DBBC rack together with the Timing board.

The Timing board receiving the external 1PPS, produces two copies, one for the PCI 9111HR board, one for the FiLa first board. Internally to the Timing board five differential high resolution 1PPS are generated, and four of them are routed to the four ADB1. These signals are used to synchronize the sampling in the four Ifs. The synchronization is performed with 1PPS pulses lasting one period of sampling clock, so in the standard case, the pulses are long 1/1.024x10^9 = 0.9765625 ns.

An additional small board making part of this set is the 1PPS-PCI adapter used to produce from a standard 1PPS, as normally present in a station, long like one clock cycle of 5 MHZ (200ns), a 280 microsecond long pulse for software synchronization. The small board is attached on the PCI9111HR connector and support also the connection of the four analog levels coming from the Conditioning Modules.

7.2 Maintenance and calibration

No maintenance or calibration is required.

8. FiLa Boards

8.1 Description

The FiLa board is used is the first and last board in the DBBC Stack, with a number of different elements populating the board in the different positions. In both positions the FiLa boards represent the digital interface between the extern and the DBBC Stack.

The FiLa in the initial position perform the following tasks:

a) data I/O interface

b) JTAG interface

c) 1PPS input

The FiLa in the final position performs the following tasks:

a)  double VSI interface

b)  DA monitor

The first board is connected through two flat cables with the PCI 7200 that operates at TTL levels, while the DBBC Stack can accept a maximum of 3.3V. So a level conversion is applied to data from the PCI7200, while data to the PCI board can be accepted with the LVTTL levels so no conversion is performed. A small connector join the Xilinx download cable to the DBBC Stack trough the FiLa and the JTAG Boundary Scan protocols can be adopted for configuration downloads and the scanning of devices.

Termination of the JTAG chain is realized in the last FiLa with a jumper. Whether additional element should be added to the JTAG chain, the jumper should be removed and the chain closed in the last element. The chain could be continued connecting from the last FiLa JTAG connector.

The last FiLa support the functionality of two independent VSI interfaces as DOM, with the two VSI connectors to be connected with a dedicated cable just on the board. In case of critical or long connections an equalizer scheme could be activated. For normal length, up to 2 m cables, no equalization is normally necessary.

A digital to analog converter is used to produce analog representation of a digital 12-bit monitor data. Not all the FPGA configurations support such functionality that could be very useful for debugging purposes. The analog output from the DA converter can be single-ended or differential and is routed to external Service panel.

8.2 Maintenance and calibration

No maintenance or calibration is necessary.

9. ADB1 Board

9.1 Description

The ADB1 Board is the analog interface of the DBBC Stack. It’s able to produce a 8-bit representation of the input analog signal, that can be single-ended or differential. The board presents four connectors in the top side, two for differential clock, two for differential input signal to be sampled. In the normal conditions single-end inputs are adopted, and the negative unused input should be terminated with a SMA 50 ohm load.

The front edge of the board has two additional connectors used for a differential reset coming from the Timing board. This signal is used to synchronize all the ADB1 boards of the stack.

The sampling clock is 1.024 GHz and is routed to the MAX108 device that produces a digital 8-bit representation of the analog input signal. The input band is pre-selected by the Conditioning Module in the ranges 10-512 MHz and 512 – 1024 MHz so that any ambiguity is avoided. In order to slow-down the clock frequency to properly insert data in the HIS bus a demultiplexed representation in two bus is produced and the data represented in DDR. Clock propagated through the HIS bus is then 256 MHz so as the maximum spectral contend in the data path. Logic level is LVPECL operating at 3.3V. For a general use of the ADB1 board several jumpers are present, but no modification is generally necessary with respect to the configuration set at the factory.

8.2 Maintenance and calibration

No maintenance or calibration is necessary.

10. Core2 Board

10.1 Description

The Core2 board is the processing element of the data sampled by the ADB1. A configuration file is loaded to the programmable device through a JTAG connection and when the device is programmed a green led placed on the front edge will light. There are two groups of  leds: sixteen green leds from the top side and four in the lower side. The meaning of the first group of  leds is shown in the table 10.1.

| 1 | DAC monitor clock |
| --- | --- |
| 2 | DAC enable |
| 3 | BBC0 overflow |
| 4 | BBC1 overflow |
| 5 | BBC2 overflow |
| 6 | BBC3 overflow |
| 7 | 10KHz retuning enable |
| 8 | 80 Hz continuous cal enable |
| 9 | Bwd code bit2 |
| 10 | Bwd code bit1 |
| 11 | Bwd code bit0 |
| 12 | 1PPS sync enable |
| 13 | Board Write Enable |
| 14 | DCM2 Lock Status |
| 15 | DCM1 Lock Status |
| 16 | 1PPS |

Table 10.1. Core1 Led indicators, upper group, all green

NOTE:

bwd code

000  		32MHz		  	@ 64 MHz clk (not yet implemented)

001  		16 MHz		@ 32 MHz clk

010  	    	8 MHz			@ 32 MHz clk

011  		4 MHz			@ 32 MHz clk

100  		2 MHz			@ 32 MHz clk

101  	    	1 MHz			@ 32 MHz clk

110  	       0,50 MHz			@ 32 MHz clk (not yet implemented)

111  	       0,25 MHz			@ 32 MHz clk (not yet implemented)

The definition of the second group of  leds is indicated in the table 10.2

| 1 | Device configured, green |
| --- | --- |
| 2 | 2.5V Vcco, yellow |
| 3 | 1.0V Vcore, yellow |
| 4 | 2.5V Vaux, yellow |

Table 10.2. Core1 Led indicators, lower group

In a normal working state it should be found on:  the leds indicating the bandwidth of the down converter, both leds indicating the correct lock status in the DCM1 and DCM2, 1PPS led pulsing at the rate of one pulse per second, the configuration of the device, yellow leds indicating the correct status of the voltage rails.

Data are available on the HIS bus and after transit to the FPGA are available to the HSIR bus for the following elements in the chain. Data converted are available on the HSO bus, and the proper cascade of the stack is assured with the HSOR bus. All the auxiliary signals for communication, monitoring, service, programming re realized in the third bus, the CCM/CCMR.

The FPGA device requires an important cooling so a heat sink cooler is placed on the top side of the device. Air flow is guaranteed by the lower side of the stack. No obstruction should be placed in the normal air path.

10.2 Maintenance and calibration

In normal conditions no maintenance or calibration is necessary.

11. Connection Set

11.1 Description

The set includes all the cabling necessary to route the signals from one board to the other or between input panels or output VSI connections. Everything is included and distinction can be done between RF connections and digital cables. The input signals from the receiver are directly routed without cabling to the Conditioning Modules, so that no cable is involved. The rear connector panel is instead connected to the different sections using small SMA pre-cabled cables, so a all the clock and analog signal distribution to the ADB1 boards. Digital cables include the connections between the PCI boards and the FiLa first and the Conditioning Modules. VSI output is using two cables directly joining the internal FiLa last to the user device (e.g. MK5B/B+).

11.2 Maintenance and calibration

No maintenance or calibration is required.

12. Firmware Set

12.1 Description

A number of configuration files is necessary for the proper working status of the system. Each Core board has its own set with a different file for each bandwidth supported. At the time of system configuration a file contains the boards present on the DBBC Stack and the file to be downloaded for the particular board in the stack. A modification in any Core modification requires the reconfiguration of the entire set of boards. This takes about 30 sec and it’s realised at start-up time and after a system command requiring a modification in one or more Core boards.

12.2 List of configurations files

For any board, indicated with Y, it’s at present available the following list of files. The actual file name will show the Core personality instead of  Y, that ranges between 1 and 16.

dbbc2_4ch_U.bit    (tunable, all the bands)

dbbc2_poly.bit       (polyphase filter)

The files are placed in the directory c:\DBBC_conf\FilesDBBC

Additional configuration files will be added at the time of their development completion.

13. Software Set

13.1 Description

The DBBC Software Set is at present running under Windows XP due to some incompatibility between drivers operating in the system under different version of Linux. As soon as it will be available a common Linux release for all the involved drivers, it will be possible to adopt Linux operating system for those stations that would require this.

The main program to be used for standard operations with the DBBC is "DBBC Control". Running this software it will become available a console able to accept commands and to show output response. No aditional software is to be run, otherwise a conflict between different programs could occur in common resources. A socket is available for having the possibility to run the system under a remote control, as it could be the Field System. At presnt is planned to develop a class of FS station commands able to communicate through this socket with the “dbbc control” program.

A file named "dbbc_config_file.txt" (in the directory c:\dbbc_conf) is to be configured in advance to run the control software to take care of the status of presence of boards in the DBBC Stack, and to set the initial conditions of the system in terms of bandwidth and frequency settings. Any additional modification in a single bbc will modify the status of the control file, that so will maintain the last settings for a further reinizialization of the system. A typical content for this file is the following:

1 dbbc2.bit  100.00 4

1 dbbc2.bit  140.99 4

1 dbbc2.bit  170.99 4

1 dbbc2.bit  230.99 4

1 dbbc2.bit  340.99 4

1 dbbc2.bit  420.99 4

1 dbbc2.bit  470.99 4

1 dbbc2.bit  490.99 4

1 dbbc2_.bit  192.99 4

1 dbbc2.bit 207.99 4

1 dbbc2.bit 217.99 4

1 dbbc2.bit 247.99 4

1 dbbc2.bit 267.99 4

1 dbbc2.bit 272.99 4

1 dbbc2.bit 100.99 4

1 dbbc2.bit 100.99 4

each row representing one bbc. The first parameter can be 0 or 1 depending on the presence of a Core2 at an adddress between 1 and 16, the second parameter is the configuration file to be loaded, the third the frequency setting, the fourth the bandwidth. 

Additional programs available only for test reasons and not to be run while "dbbc Control" is taking the control of the system are:

1) “agc_if “    it is used to manually set and monitor the status of the Conditioning Modules present in the system.

2) “ad9858”      used to set the system clock to 1024MHz.

3)  “test_poly16” to run the polyphase filter configuration

14. DBBC Command Set

Description

This chapter describes the basic commands the DBBC is able to recognize with the control console. The structure and the meaning of the different commands is Field System based, so to simplify the dialogue with the FS and minimize efforts on the FS side. Any commands sent to the interpreter from the DBBC console is then identical to the command sent from the Field System environment. Similarly output information issued by any command are reported in FS style.

List of commands

At present 10 commands are defined for the main functionalities:

1)

DBBCnn = freq, IF, bwdU, bwdL, gainU, gainL, tpint  	      (DownConverterConfiguration)

where

nn => 01, .. ,16   indicates the number of CoreModule;

freq => is the base band frequency in MHz, in the range 0010.000000 - 2,048.000000;

IF =>  A or B or C or D. Any Core2 is connected to a band in the standard communication so this value is only informative.

bwdU => band width of the upper side, in MHz;

bwdL => band width of the lower side, in MHz;

bwdL and bwdU are always the same in a single bbc.

gainU => gain of the upper side in the range 0 - 255, step 1

gainL => gain of the upper side in the range 0 - 255, step 1

this values could normally kept to 1 as the magnitude bit is controlled with a dinamic threshold.

tpint => total power integration time in seconds, in the range 1 - 60, step 1

(1 sec implemented at present).

DBBCnn

reports the setting of the CoreModule nn, including cal tone off.

DBBCnn/freq,IF,bwdU,bwdL,gainU,gainL,tpU/calon,tpL/calon,tpUcaloff,tpLcaloff

2)

DBBCIF(A,B,C,D) =  input_ch, gain, filter

where

input => input channel of the four possible (1,2,3,4) .

gain => the gain of the channel is set in manual mode if a number is indicated in the range  -16 to +16 dB, step 0.5 dB. If AGC is indicated the gain is automatically set so to satisfy the optimal level for the analog to digital converter.

filter=> 2 (10-512 MHz), 1 (512-1024 MHz), 3 (ext 1), 4 (ext 2)

DBBCIF

reports the settings of the IFs modules.

3)

DBBCFORM = VSI1 mode, VSI2 mode

where

VSI1/2mode => is the mapping of the 64 channels in the VSI1/2 interface. Possible predefined values are: GEO, ASTRO.

DBBCFORM

reports the settings of the VSI output mapping.

4)

DBBCMON= bnn[u|l]

set the Digital to Analog Channel source.

nn  =>  01,2,3   indicates the number of  band;

u/l => upper or lower side band

DBBCMON

reports the Digital to Analog Channel source.

5)

DBBC_CAL_IF

reports the entire system total power and gain settings in the IF units. The output is:

DBBC_CAL_IF=tp_ifa, gain_ifa, tp_ifb, gain_ifb, tp_ifc, gain_ifc,, tp_ifd, gain_ifd

Not yet active

6)

DBBC_CAL_CH

reports the entire system total power and gain settings in the converted channels. The output is:

DBBC_CAL_IF=tp_l1, gain_l1, tp_u1, gain_u1, tp_l2, gain_l2, tp_u2, gain_u2, ..., tp_l16, gain_l16, tp_u16, gain_u16

Not yet active

7)

PPS_SYNC

Synchronize to external 1pps

8)

DBBCGAIN=dbbcnn,gainU,gainL

Adjust  gain level

dbbcnn = 1-16

gainU = 0 - 255

gainL  = 0 -255

9)

RESETALL

System reset

10)

RECONF

System reconfiguration

15. Schematics Drawing 		In preparation

15.1 Connection Set

15.2 Power Distributor

15.3 Conditioning Module

15.4 CaT Set

15.5 FiLa

15.6 ADB1

15.7 Core1

16. Firmware _4ch register description (dbbc2.bit)

Register       		DataW                	DataR

(WD26 - WD22)  	(WD15 - WD00)   	(RD32 -RD00)    	Meaning

0x0		NA			rd15-rd00       		statistics output L

0x0		NA			rd31-rd16       		statistics output U  			0x0		wd01-wd00		 NA                		statistics  selection						  					0x0 = MU - ML

0x1 = MU -  SL

0x2 = SU  -  ML

0x3 = SU  -  SL

0x1		NA	           		rd15-rd00        	Total Power L / calon

0x1		NA	           		rd31-rd16        	Total Power U / calon

0x2		wd07-wd00        	rd07-rd00       		amplification L

0x2		wd15-wd08         	rd15-rd08       		amplification U

0x3		wd03-wd00         	rd03-rd00       		DA monitor selection

0x0 = na

0x1 = na

0x2 = U

0x3 = L

0x4 = MU

0x5 = SU

0x6 = ML

0x7 = SL

0x8 = Input data phase 0

0x9 = DDS cos phase 0

0xA = Mixer out I phase 0

0xB = Poly out Q

0xC = 1:4 out Q

0xD = Q out

0xE = I out

0x4		wd-15-wd00       	rd15-rd00        	Phase Offset

0x10		wd-15-wd00       	rd15-rd00       		VSI1 high output selection

NA

0x11		wd-15-wd00       	rd15-rd00       		VSI1 low output selection										Geo track assignment= 0

Astro track assignment=1

Pattern = 2

0x12		wd-15-wd00       	rd15-rd00       		VSI2 high output selection

NA

0x13		wd-15-wd00       	rd15-rd00       		VSI1 low output selection

Fixed = 0|1

Pattern = 2

0x14		wd-15-wd00       	rd15-rd00        	Phase increment high

0x15		wd-15-wd00       	rd15-rd00        	Phase increment low

0x16		Any               		NA      	         		Phase increment load

0x23		wd15-wd0       	rd15-rd0       		General control

1pps out selection

bit0 	0 = 1pps 1us ??

1 = 1pps 32  ??

Total power mode

bit1 	0 = normal

1 = cont. cal

10KHz DDS reset

bit2 	0 = off

1 = on

DAC board out (only write)

bit3	0 = off

1= on

DCM256 status (only read)

bit3 	0 = unlock

1 = lock

DAC clock inv (only write)

bit4 	0 = inverted

1 = not inverted

DCM128 status (only read)

bit4 	0 = unlock

1 = lock

lPPS tuning load (NA)

bit5 	0 = disabled

1 = enabled

bit6	NA

bit7	NA

bit8	NA

bit9	NA

DAC Monitor clock set

bit10	0 = 128 MHz

1 = 32 MHz

1PPS sync

bit11	0 = no ext 1pps sync

1 = ext 1pps sync

0x24		NA		 	rd15-rd00        	Total Power L / caloff

0x24		NA		 	rd31-rd16        	Total Power U / caloff

0x28		wd02-wd00    		rd02-rd00		Band Selection

0 = 32 MHz  (NA)

1 = 16 MHz

2 =  8 MHz

3 =  4 MHz

4 =  2 MHz

5 =  1 MHz

6 =  05 MHz  (NA)

7 =  025MHz (NA)

0x29		wd11-wd00		rd11-rd00		Magnitude threshold U

0x30		wd11-wd00		rd11-rd00		Magnitude threshold L

0x3E		Any       		NA      	       		DDS reset

0x3F		Any       		NA      	       		DCM reset