{"text_format":"2","workbook":"xls","n_sheets":3,"sheets":["Sheet1","Sheet2","Sheet3"],"not_rendered":{"charts":4,"images":0,"macros":0,"external_links":1,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"run_cost":"one","rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900","formulas":"cached_only"}

{"sheet":"Sheet1","index":0,"n_rows":47,"n_cols":19,"rows_kept":47,"cut":false,"hidden":false,"merged":[[8,2,8,3],[8,6,8,7],[16,1,21,1]],"flags":["formulas_cached_only"]}
["Linearity tests 9/6/2011"]
[]
["Tests for linearity of power measurements from various instruments",{"$nulls":12},100]
["Connected signal generator SML01 (SN#     ) to input of DUT"]
["SML Freq. = 300 MHz for 8757D + FSU testing, 620 MHz for DBBC to be in Filter 1 band!"]
["SML connected to DUT by 5m MF141 Cable, loss = -2.64 dB@300 MHz as measured by 8757D"]
[]
[]
["Instrument",null,"8757D Scalar Net. Analyser",null,null,null,"FSU 26 Spec. An.",null,null,null,"DBBC"]
[null,null,"SML P dBm","Analyser P dBm","Corrected An. P dBm",null,"SML P dBm","FSU26 P dBm","Corrected P",null,"SML P dBm","DBBC P [relative)",{"$nulls":5},"DBBC Counts","P(W)"]
[null,null,0,-2.2,0.43999999999999995,null,0,-2.12,0.52,null,0,62742,null,null,62742,0,null,50000,9.174871702976619e-05]
[null,null,-5,-7.223,-4.583,null,-5,-7.11,-4.470000000000001,null,-5,56550,null,6192,56550,-5]
[null,null,-10,-12.21,-9.57,null,-10,-12.11,-9.469999999999999,null,-10,49326,null,7224,49326,-10]
[null,null,-15,-17.16,-14.52,null,-15,-17.09,-14.45,null,-15,42554,null,6772,42554,-15]
[null,null,-20,-22.12,-19.48,null,-20,-22.1,-19.46,null,-20,35302,null,7252,35302,-20]
[null,null,-25,-26.95,-24.31,null,-25,-27.09,-24.45,null,-25,27751,null,7551,27751,-25]
[null,"noise in reading of 8757D. Turned averaging on",-30,-31.5,-28.86,null,-30,-32.1,-29.46,null,-30,20801,null,6950,20801,-30]
[null,null,-35,-35.3,-32.66,null,-35,-37.07,-34.43,null,-35,14151,null,6650,14151,-35]
[null,null,-40,-37.85,-35.21,null,-40,-42.08,-39.44,null,-40,6960,null,7191,6960,-40]
[null,null,-45,-39.1,-36.46,null,-45,-47.01,-44.37,null,-45,1669,null,null,1669,-45]
[null,null,-50,-39.61,-36.97,null,-50,-52.09,-49.45,null,-50,0,null,6972.75]
[null,null,-55,-39.8,-37.16,null,-55,-57.07,-54.43,null,-55,0]
{"$rows_empty":24}
[null,null,"8757 in DC mode."]

{"sheet":"Sheet2","index":1,"n_rows":0,"n_cols":0,"rows_kept":0,"cut":false,"hidden":false,"merged":[],"flags":["formulas_cached_only"]}

{"sheet":"Sheet3","index":2,"n_rows":0,"n_cols":0,"rows_kept":0,"cut":false,"hidden":false,"merged":[],"flags":["formulas_cached_only"]}