{"text_format":"2","workbook":"xls","n_sheets":3,"sheets":["12m_06seb10","12m_28feb11","12m_03Mar11"],"not_rendered":{"charts":0,"images":0,"macros":0,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"run_cost":"one","rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900","formulas":"cached_only"}

{"sheet":"12m_06seb10","index":0,"n_rows":10,"n_cols":9,"rows_kept":10,"cut":false,"hidden":false,"merged":[],"flags":["formulas_cached_only"]}
[]
["2006-09-01"," Jim Lovell"]
[]
["System temperature calculations assume a Load temperature of 290K, LNA temperature of 45K and Sky temperature of 5 K"]
["Y-factors"]
["Band (and freq)","Y LCP","Y RCP","Tsys (LCP)","Tsys (RCP)","Diode (cal units) LCP","Diode (cal units) RCP","Diode (K) LCP","Diode (K) RCP"]
["S 2260",5.88,5.48,86.50571638625897,94.85163185927917,29,25.5,2.982955737457206,3.7196718376187907]
["S 2300",5.99,6.02,84.34217707992167,83.76156961683044]
["S 2350",6.24,5.71,79.62414960193334,89.95903892956]
["X broadband",5.997,6.18,84.20634305086415,80.73183186100091,13.6,12.1,6.19164287138707,6.672052219917431]

{"sheet":"12m_28feb11","index":1,"n_rows":24,"n_cols":7,"rows_kept":24,"cut":false,"hidden":false,"merged":[],"flags":["formulas_cached_only"]}
["Observations by Stas and Jamie, 28 Feb 2011"]
[]
["T_load",293,"K"]
["T_LNA",45,"K"]
[]
["No cal"]
["Band/polarization","P(LNA+load)/dB","P(sys)/dB","dP","Y","T_sys/K"]
["S,LCP",-19.566,-24.46,4.894000000000002,3.0860289794360662,109.52586714262331]
["S,RCP",-22.756,-27.504,4.748000000000001,2.9840081159302208,113.2704694050852]
["X,LCP",-7.7743,-13.537,5.762700000000001,3.7693806778762777,89.66990306493372]
["X,RCP",-7.6729,-13.455,5.7821,3.7862562218012297,89.27023957169064]
[]
["With cal"]
["Band/polarization","P(LNA+load+cal)/dB","P(sys+cal)/dB","dP","Y","T_cal/K","T_cal,2/K"]
["S,LCP",-19.484,-24.215,4.730999999999998,2.9723503619120284,6.312645765342079,6.442474265246631]
["S,RCP",-22.681,-27.29,4.608999999999998,2.8900143553494697,5.633141013980426,5.887745769044386]
["X,LCP",-7.6238,-12.999,5.3752,3.447624831066605,11.78767890723933,11.91833585586329]
["X,RCP",-7.5388,-12.964,5.4252,3.4875464451032303,10.719756964027265,10.599452173490647]
[]
["Diode height (sky)"]
["S,LCP",0.245000000000001]
["S,RCP",0.2140000000000022]
["X,LCP",0.5380000000000003]
["X,RCP",0.49099999999999966]

{"sheet":"12m_03Mar11","index":2,"n_rows":83,"n_cols":7,"rows_kept":83,"cut":false,"hidden":false,"merged":[],"flags":["formulas_cached_only"]}
["March 3 2011",null,"Ambient temp: 13C"]
["Observations by Jim Lovell"]
["T_load",286,"K"]
["T_LNA",45,"K"]
[]
["Summary of results"]
["No cal"]
["Band/polarization","P(LNA+load)/dB","P(sys)/dB","dP","Y","T_sys/K"]
["S,LCP",-19.0964,-23.898,4.8016000000000005,3.0210645167041585,109.56402889439305]
["S,RCP",-22.413,-27.133,4.719999999999999,2.9648313895243423,111.6420991660856]
["X,LCP",-7.214,-13.086,5.872,3.865449465982839,85.63040415167842]
["X,RCP",-7.0546,-12.956,5.9014,3.891705785478857,85.05267824588952]
[]
["With cal"]
["Band/polarization","P(LNA+load+cal)/dB","P(sys+cal)/dB","dP","Y","T_cal/K","T_cal,2/K"]
["S,LCP",-19.0096,-23.645,4.635400000000001,2.907635750261194,6.514719931365684,6.682064078687802]
["S,RCP",-22.334,-26.893,4.559000000000001,2.856932634583698,6.487065429641256,6.0761258632487625]
["X,LCP",-7,-12.297,5.297000000000001,3.386101717130323,17.2024273603908,16.718657162291834]
["X,RCP",-6.858,-12.226,5.368000000000001,3.441913882684293,15.6664025939929,15.32831010752966]
[]
["Diode height (sky)"]
["S,LCP",0.2530000000000001]
["S,RCP",0.23999999999999844]
["X,LCP",0.7889999999999997]
["X,RCP",0.7299999999999986]
[]
[]
["Raw data"]
["Switch pos"]
["Data sequence",null,"Level (dB)"]
["Sky S-band LCP","T",-23.899]
["diode on","P",-23.644]
["diode off","Q"]
["Sky S-band RCP","U",-27.135]
["diode on","P",-26.892]
["diode off","Q"]
["Sky X-band LCP","V",-13.09]
["diode on","P",-12.303]
["diode off","Q"]
["Sky X-band RCP","W",-12.948]
["diode on","P",-12.218]
["diode off","Q"]
["Sky S-band LCP","T",-23.898]
["diode on","P",-23.645]
["diode off","Q"]
["Sky S-band RCP","U",-27.133]
["diode on","P",-26.893]
["diode off","Q"]
["Sky X-band LCP","V",-13.086]
["diode on","P",-12.297]
["diode off","Q"]
["Sky X-band RCP","W",-12.956]
["diode on","P",-12.226]
["diode off","Q"]
[]
[]
["LOAD ON"]
["Sky X-band RCP","W",-7.0546]
["diode on","P",-6.858]
["diode off","Q"]
["Sky X-band LCP","V",-7.214]
["diode on","P",-7]
["diode off","Q"]
["Sky S-band LCP","T",-19.0964]
["diode on","P",-19.0096]
["diode off","Q"]
["Sky S-band RCP","U",-22.413]
["diode on","P",-22.334]
["diode off","Q"]
["LOAD OFF"]
[]
["Sky S-band LCP","T",-23.925]
["diode on","P",-23.667]
["diode off","Q"]
["Sky S-band RCP","U",-27.1388]
["diode on","P",-26.898]
["diode off","Q"]
["Sky X-band LCP","V",-13.1009]
["diode on","P",-12.3124]
["diode off","Q"]
["Sky X-band RCP","W",-12.9851]
["diode on","P",-12.2543]
["diode off","Q"]