{"text_format":"2","workbook":"xls","n_sheets":4,"sheets":["12m_06seb10","12m_28feb11","12m_03Mar11","12mhb 07Mar11"],"not_rendered":{"charts":0,"images":0,"macros":0,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"run_cost":"one","rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900","formulas":"cached_only"}

{"sheet":"12m_06seb10","index":0,"n_rows":10,"n_cols":9,"rows_kept":10,"cut":false,"hidden":false,"merged":[],"flags":["formulas_cached_only"]}
[]
["2006-09-01"," Jim Lovell"]
[]
["System temperature calculations assume a Load temperature of 290K, LNA temperature of 45K and Sky temperature of 5 K"]
["Y-factors"]
["Band (and freq)","Y LCP","Y RCP","Tsys (LCP)","Tsys (RCP)","Diode (cal units) LCP","Diode (cal units) RCP","Diode (K) LCP","Diode (K) RCP"]
["S 2260",5.88,5.48,86.50571638625898,94.8516318592792,29,25.5,2.9829557374572064,3.719671837618792]
["S 2300",5.99,6.02,84.34217707992167,83.76156961683046]
["S 2350",6.24,5.71,79.62414960193335,89.95903892956]
["X broadband",5.997,6.18,84.20634305086415,80.73183186100093,13.6,12.1,6.19164287138707,6.672052219917433]

{"sheet":"12m_28feb11","index":1,"n_rows":24,"n_cols":7,"rows_kept":24,"cut":false,"hidden":false,"merged":[],"flags":["formulas_cached_only"]}
["Observations by Stas and Jamie, 28 Feb 2011"]
[]
["T_load",293,"K"]
["T_LNA",45,"K"]
[]
["No cal"]
["Band/polarization","P(LNA+load)/dB","P(sys)/dB","dP","Y","T_sys/K"]
["S,LCP",-19.566,-24.46,4.894000000000002,3.086028979436066,109.52586714262333]
["S,RCP",-22.756,-27.504,4.748000000000001,2.9840081159302203,113.27046940508522]
["X,LCP",-7.7743,-13.537,5.762700000000001,3.7693806778762773,89.66990306493373]
["X,RCP",-7.6729,-13.455,5.7821,3.7862562218012297,89.27023957169064]
[]
["With cal"]
["Band/polarization","P(LNA+load+cal)/dB","P(sys+cal)/dB","dP","Y","T_cal/K","T_cal,2/K"]
["S,LCP",-19.484,-24.215,4.730999999999998,2.972350361912028,6.31264576534208,6.442474265246631]
["S,RCP",-22.681,-27.29,4.608999999999998,2.8900143553494693,5.633141013980428,5.887745769044443]
["X,LCP",-7.6238,-12.999,5.3752,3.4476248310666047,11.787678907239309,11.918335855863347]
["X,RCP",-7.5388,-12.964,5.4252,3.48754644510323,10.719756964027267,10.59945217349059]
[]
["Diode height (sky)"]
["S,LCP",0.245000000000001]
["S,RCP",0.2140000000000022]
["X,LCP",0.5380000000000003]
["X,RCP",0.49099999999999966]

{"sheet":"12m_03Mar11","index":2,"n_rows":83,"n_cols":7,"rows_kept":83,"cut":false,"hidden":false,"merged":[],"flags":["formulas_cached_only"]}
["March 3 2011",null,"Ambient temp: 13C"]
["Observations by Jim Lovell"]
["T_load",286,"K"]
["T_LNA",45,"K"]
[]
["Summary of results"]
["No cal"]
["Band/polarization","P(LNA+load)/dB","P(sys)/dB","dP","Y","T_sys/K"]
["S,LCP",-19.0964,-23.898,4.8016000000000005,3.021064516704158,109.56402889439306]
["S,RCP",-22.413,-27.133,4.719999999999999,2.964831389524342,111.64209916608561]
["X,LCP",-7.214,-13.086,5.872,3.865449465982839,85.63040415167842]
["X,RCP",-7.0546,-12.956,5.9014,3.891705785478857,85.05267824588952]
[]
["With cal"]
["Band/polarization","P(LNA+load+cal)/dB","P(sys+cal)/dB","dP","Y","T_cal/K","T_cal,2/K"]
["S,LCP",-19.0096,-23.645,4.635400000000001,2.9076357502611936,6.514719931365685,6.682064078687802]
["S,RCP",-22.334,-26.893,4.559000000000001,2.8569326345836976,6.487065429641257,6.076125863248819]
["X,LCP",-7,-12.297,5.297000000000001,3.3861017171303227,17.202427360390825,16.718657162291777]
["X,RCP",-6.858,-12.226,5.368000000000001,3.4419138826842923,15.66640259399293,15.328310107529603]
[]
["Diode height (sky)"]
["S,LCP",0.2530000000000001]
["S,RCP",0.23999999999999844]
["X,LCP",0.7889999999999997]
["X,RCP",0.7299999999999986]
[]
[]
["Raw data"]
["Switch pos"]
["Data sequence",null,"Level (dB)"]
["Sky S-band LCP","T",-23.899]
["diode on","P",-23.644]
["diode off","Q"]
["Sky S-band RCP","U",-27.135]
["diode on","P",-26.892]
["diode off","Q"]
["Sky X-band LCP","V",-13.09]
["diode on","P",-12.303]
["diode off","Q"]
["Sky X-band RCP","W",-12.948]
["diode on","P",-12.218]
["diode off","Q"]
["Sky S-band LCP","T",-23.898]
["diode on","P",-23.645]
["diode off","Q"]
["Sky S-band RCP","U",-27.133]
["diode on","P",-26.893]
["diode off","Q"]
["Sky X-band LCP","V",-13.086]
["diode on","P",-12.297]
["diode off","Q"]
["Sky X-band RCP","W",-12.956]
["diode on","P",-12.226]
["diode off","Q"]
[]
[]
["LOAD ON"]
["Sky X-band RCP","W",-7.0546]
["diode on","P",-6.858]
["diode off","Q"]
["Sky X-band LCP","V",-7.214]
["diode on","P",-7]
["diode off","Q"]
["Sky S-band LCP","T",-19.0964]
["diode on","P",-19.0096]
["diode off","Q"]
["Sky S-band RCP","U",-22.413]
["diode on","P",-22.334]
["diode off","Q"]
["LOAD OFF"]
[]
["Sky S-band LCP","T",-23.925]
["diode on","P",-23.667]
["diode off","Q"]
["Sky S-band RCP","U",-27.1388]
["diode on","P",-26.898]
["diode off","Q"]
["Sky X-band LCP","V",-13.1009]
["diode on","P",-12.3124]
["diode off","Q"]
["Sky X-band RCP","W",-12.9851]
["diode on","P",-12.2543]
["diode off","Q"]

{"sheet":"12mhb 07Mar11","index":3,"n_rows":12,"n_cols":6,"rows_kept":12,"cut":false,"hidden":false,"merged":[],"flags":["formulas_cached_only"]}
["2011-03-07",null,"Ambient temp: 20C"]
["Observations by Jmc & Jim Lovell"]
["T_load",293,"K"]
["T_LNA",45,"K"]
[]
["Summary of results"]
["No cal"]
["Band/polarization","P(LNA+load)/dB","P(sys)/dB","dP","Y","T_sys/K"]
["S,LCP",-19.16,-24.66,5.5,3.548133892335755,95.26134307673853]
["S,RCP",-22.33,-27.56,5.23,3.3342641276323497,101.3716931417826]
["X,LCP",-7.44,-13.53,6.089999999999999,4.0644332916521275,83.1604250201898]
["X,RCP",-7.68,-13.75,6.07,4.045758916974427,83.5442761015452]